## Slide 1
Time and Frequency
Calibration of
Stopwatches
Engineer Liz Catherine Hernández Forero
Engineer Eduin Culma Caviedes

## Slide 2
DIRECT TIME MEASUREMENT
Reference Clock
Reference Clock
Reference Clock
Reference Clock
Clock
under calibration
Clock
under calibration
Clock
under calibration
Clock
under calibration
Time

## Slide 3
INDIRECT FREQUENCY MEASUREMENT
Metal Shielding
Conductive surface
Amplifier stage
Amplifier control
Frequency counter
Oscilloscope
Typical signal

## Slide 4
INDIRECT FREQUENCY MEASUREMENT

## Slide 5
http://www.jewelerssupplies.com/product7141.html
INDIRECT FREQUENCY MEASUREMENT

## Slide 6
DIRECT TIME MEASUREMENT
Reference Clock
Reference Clock
Reference Clock
Reference Clock
Clock
under calibration
Clock
under calibration
Clock
under calibration
Clock
under calibration
Time

## Slide 7
Metrological confirmation
Stopwatch
Start / stop mechanism (20 ms)
Stopwatch operation time BIG
Small
Start / stop
DIFFERENCES OF
ACCUMULATED TIME
OF MEASURING
versus

## Slide 8
Metrological confirmation
Stopwatch
versus
Small
Start / stop
DIFFERENCES OF
ACCUMULATED TIME
OF MEASURING
Linear adjustment of measurements
Time differences
Time

## Slide 9
Stopwatch
Linear adjustment of measurements
Time differences
Time
Time differences
Time

## Slide 10
T = 24 h
T = 86 400 s
u_operator 0.2 s
u_relative
u_relative
T = 10 min
T = 600 s
u_relative
u_relative
Linear adjustment of measurements
Time differences
Time
Time differences
Time

## Slide 11
Repeatability and Reproducibility
Fuente: http://www.caletec.com/blog/tag/estudios-rr/
¿u_operator 0.2 s?
¿u_operator 0.02 s?

## Slide 12
Reference Clock
Reference Clock
Reference Clock
Reference Clock
Clock
under
Clock
under calibration
Clock
under calibration
Clock
under calibration
Time
calibration
DIRECT TIME MEASUREMENT

## Slide 13
Fuente:
Introduction to Frequency Measurements (Generación y medición de frecuencia).
Francisco Jiménez, CENAM.
SIM TFMWG Workshop and Planning Meeting at CENAMEP
Panama City, Panama (January 27th through January 29th, 2015)
STABILITY
Allan Deviation
t is the time measured by the standard.
tstopwatch is the time measured by the instrument under calibration.
2 is the coverage factor k.
ACCURACY
Frequency
Offset
DIRECT TIME MEASUREMENT

## Slide 14
Characteristic | Year 2014 | 
 | SEIKO S025-6000 | CASIO HS-3
Data example | -0.09 s -0.13 s -0.15 s | 0.04 s 0.01 s -0.02 s
Analysis (Advance / Delay) | Advance 0.0288 s (0.03 s) for each hour of continuous use. | Advance 0.0324 s (0.03 s) for each hour of continuous use.
Characteristic | Year 2014 | 
 | SEIKO S025-6000 | CASIO HS-3
 | -0.09 s -0.13 s -0.15 s | 0.04 s 0.01 s -0.02 s
Analysis (Advance / Delay) | Advance 0.0288 s (0.03 s) for each hour of continuous use. | Advance 0.0324 s (0.03 s) for each hour of continuous use.
INDIRECT FREQUENCY MEASUREMENT

## Slide 15
INDIRECT FREQUENCY MEASUREMENT
Metal Shielding
Conductive surface
Amplifier stage
Amplifier control
Frequency counter
Oscilloscope
Typical signal

## Slide 16
F_stopwatch = {f_quartz = 32.0 Hz}
F_stopwatch = {f_quartz / (2^10)}
F_stopwatch = {f_quartz = 32 768 Hz / (1024)}

## Slide 17
Fuente:
Introduction to Frequency Measurements (Generación y medición de frecuencia).
Francisco Jiménez, CENAM.
SIM TFMWG Workshop and Planning Meeting at CENAMEP
Panama City, Panama (January 27th through January 29th, 2015)
t is the time measured by the standard.
tstopwatch is the time measured by the instrument under calibration.
2 is the coverage factor k.
STABILITY
Allan Deviation
ACCURACY
Average
Frequency
Offset
INDIRECT FREQUENCY MEASUREMENT

## Slide 18
Characteristic | Year 2013 |  |  | 
 | EXTECH | CASIO HS-3 | COLE PARMER | Q&Q
Data example | 32.0001259079 32.0001251573 32.0001235423 | 42.6672213965 42.6672337584 42.6672332012 | 32.0007750885 32.0007718743 32.0007693736 | 128.0002218460 128.0003000120 128.0000791540
Divider |  |  |  | 
Analysis (Advance / Delay) | Advance 1 s in 71.40817 hours | Advance 1 s in 22.22223 hours | Advance 1 s in 11.97319 hours | Advance 1 s in 175.69752 hours
Characteristic | Year 2013 |  |  | 
 | EXTECH | CASIO HS-3 | COLE PARMER | Q&Q
 | 32.0001259079 32.0001251573 32.0001235423 | 42.6672213965 42.6672337584 42.6672332012 | 32.0007750885 32.0007718743 32.0007693736 | 128.0002218460 128.0003000120 128.0000791540
Analysis (Advance / Delay) | Advance 1 s in 71.40817 hours | Advance 1 s in 22.22223 hours | Advance 1 s in 11.97319 hours | Advance 1 s in 175.69752 hours
INDIRECT FREQUENCY MEASUREMENT

## Slide 19
COMPARISON
Characteristic | Year 2014 |  |  | 
 | COLE PARMER | CASIO HS-10W | SEIKO S025-6000 | CASIO HS-3
Data example | 31.99964803460 31.99963118190 31.99961866410 | 28.44442438467 28.44442408857 28.44442390715 | -0.09 s -0.13 s -0.15 s | 0.04 s 0.01 s -0.02 s
Divider |  |  | Does not apply. | Does not apply.
Analysis (Advance / Delay) | Delay 1 s in 23.86408 hours | Advance 1 s in 38.68772 hours | Advance 0.0288 s (0.03 s) for each hour of continuous use. | Advance 0.0324 s (0.03 s) for each hour of continuous use.
Characteristic | Year 2014 |  |  | 
 | COLE PARMER | CASIO HS-10W | SEIKO S025-6000 | CASIO HS-3
 | 31.99964803460 31.99963118190 31.99961866410 | 28.44442438467 28.44442408857 28.44442390715 | -0.09 s -0.13 s -0.15 s | 0.04 s 0.01 s -0.02 s
 |  |  | Does not apply. | Does not apply.
Analysis (Advance / Delay) | Delay 1 s in 23.86408 hours | Advance 1 s in 38.68772 hours | Advance 0.0288 s (0.03 s) for each hour of continuous use. | Advance 0.0324 s (0.03 s) for each hour of continuous use.

## Slide 20
¡Gracias!
Thank you!
Obrigado!
http://aprendelenguadesignos.com/signos-de-gracias-y-hola/