## Slide 1
GPS calibrations
for UTC contributors
Stefania Römisch, Bijunath Patla

## Slide 2: OUTLINE
Time transfer using GPS common-view
GNSS differential calibrations
  Definition of a local time reference point
  Local delays measurements
  Uncertainty budget
The NIST travelling system
  Description
  Setup Demo
Requirements for calibration campaigns
  Laboratory information
  Data formats
  Logistics

## Slide 3: GPS common-view
GPS receiver
GPS receiver
Clock A
Clock B
ΔtA = CLOCK A - GPStime
Clock A – Clock B = ΔtA - ΔtB
TIME TRANSFER
ΔtB = CLOCK B - GPStime
5/10 MHz
1 PPS

## Slide 4: Local Delays
GPS receiver
INTDLY
REFDLY
CABDLY
t = (Clock A +REFDLY) - (GPStime +CABDLY) + INTDLY
calibrations
Local measurements
5/10 MHz
1 PPS
Clock A –Clock B = ΔtA - ΔtB
+ (REFDLYA – REFDLYB) - (CABDLYA - CABDLYB) + INTDLYA - INTDLYB

## Slide 5: BIPM Calibration Campaigns
Based on differential calibration trips with closure.
The trip begins and ends at the same laboratory (i.e. NIST), where the reference receiver for that trip is located.
The traveling receiver is just a transfer system between the reference receiver and the visited receivers of that trip.
Calibration Identifiers are generated for every calibration trip
znnn-YYYY
1: diff. cal. with closure
2: other (i.e. manuf. cal., absolute cal. etc.)
000-009: G1
011-099: G2
101-999: other
year

## Slide 6: Differential Calibration
TRAV receiver
INTDLYT
REFDLYT,R
CABDLYT
LOCAL receiver
INTDLYR
Clock R
REFDLYR
CABDLYR
t= (GPStime +CABDLYR) − (Clock R +REFDLYR) + INTDLYR
t= (GPStime +CABDLYT) − (Clock R +REFDLYT,R) + INTDLYT
5/10 MHz
1 PPS
INTDLYR-T =
(t)R - (t)T,R - (CABDLYR - CABDLYT) + (REFDLYR – REFDLYT,R)
REFERENCE LAB
COMMON-CLOCK MEASUREMENT

## Slide 7: Local Delays: REFDLY
“REFDLY=(delay between UTC(k) and PPSin) + (delay between PPSin and RCVR’s internal ref. point)”
XP
REFDLY = XP + XO
XO
When PPSout is NOT available
OR
OR XO = 0

## Slide 8: Local Delays: REFDLY
Delay with respect to a reference point using a PIVOT
XP = PPSin – UTC(k)
SETUP
TIC: 50 Ohms, positive edge, 1V trigger point
CBL2 longer than CBL1
Multiple TIC measurements if possible
CBL2
1st measurement:
UTC(k) - PIVOT
CBL1
CBL1
CBL2
2nd measurement:
PPSin - PIVOT

## Slide 9: Local Delays: CABDLY
Delay of antenna cable
BEST: using a network analyzer at L1 and L2 center frequencies
GOOD: using a TIC and PPS signals.
CABDLY= 2nd meas – 1st meas
CBL2
CBL1
CBL2
CBL1
SETUP
TIC: 50 Ohms, positive edge, 1V trigger point
CBL2 longer than CBL1
Multiple TIC measurements if possible

## Slide 10: Differential Calibration
TRAV receiver
INTDLYT
REFDLYT,R
CABDLYT
LOCAL receiver
INTDLYR
Clock R
REFDLYR
CABDLYR
t= (GPStime +CABDLYR) − (Clock R +REFDLYR) + INTDLYR
t= (GPStime +CABDLYT) − (Clock R +REFDLYT,R) + INTDLYT
5/10 MHz
1 PPS
INTDLYR-T =
(t)R - (t)T,R - (CABDLYR - CABDLYT) + (REFDLYR – REFDLYT,R)
REFERENCE LAB
COMMON-CLOCK MEASUREMENT

## Slide 11: Differential Calibration
LOCAL receiver
INTDLYv
Clock V
REFDLYv
CABDLYv
TRAV receiver
INTDLYT
REFDLYT,V
CABDLYT
5/10 MHz
1 PPS
t= (GPStime +CABDLYV) − (Clock V +REFDLYV) + INTDLYV
t= (GPStime +CABDLYT) − (Clock V +REFDLYT,V) + INTDLYT
INTDLYV-T =
(t)V - (t)T,V - (CABDLYV - CABDLYT) + (REFDLYV - REFDLYT)
VISITED LAB
COMMON-CLOCK MEASUREMENT

## Slide 12: Differential Calibration
GPS receiver
INTDLYR
Clock R
REFDLYR
CABDLYR
5/10 MHz
1 PPS
GPS receiver
INTDLYV
Clock V
REFDLYV
CABDLYV
INTDLYV-T - INTDLYR-T = INTDLYV-R
VISITED LAB
REFERENCE LAB

## Slide 13: Differential Calibration
INTDLYk-T - INTDLYNIST-T = INTDLYLabk-NIST
G1 campaigns: INTDLYNIST-BIPM
ΔINTDLYLabk-BIPM
from “BIPM guidelines for GNSS calibration” V3.2

## Slide 14: Calibration Uncertainty
INTDLYV-R = (t)R - (t)T,R - CABDLYR + CABDLYT + REFDLYR – REFDLYT,R
-[(t)V - (t)T,V - CABDLYV + CABDLYT + REFDLYV – REFDLYT,V]
Uncertainty
of the differential calibration
Statistical analysis of RAWDIF (0.2ns)
Position error (0.05 ns) and multipath (0.3ns) (RAWDIF)
Local delay measurements (REFDLY) (0.5ns)
Antenna delay measurements (CABDLY) (0.5 ns)
Mis-closure > 2uCAL
Invalid calibration
OR
uCAL = mis-closure
Mis-closure:
INTDLYR-T, end - INTDLYR-T, beginning

## Slide 15: Calibration Uncertainty
Calibration Identifier
10nn-YYY
uCAL0
uncertainty associated to Calibration Identifier (trip)
TAI computation uses links with PTB as pivot
UTC(k)-UTC(PTB)
uCAL0│G2-PTB = 2.5 ns
Labk-NIST
NIST-BIPM
PTB-BIPM
UTC links

## Slide 16: NIST Travelling System
GPS receiver NB05
Pinwheel GPS Antenna
Antenna cable
TIC for local delay measurements
Laptop
Coax and serial cables, etc.

## Slide 17: NIST Travelling System
Inside NB05
The quartz oscillator inside the receiver is locked to the RFin, to improve its stability
PPSin enables the internal PPS generator, allowing the synchronization of the receiver’s time base to the local time reference
PPSout is the best representation of the receiver’s time base
PPSout – PPSin is continually monitored: it provides information about the receiver’s time base
GPS receiver

## Slide 18: NIST Travelling System
5/10 MHz
1 PPS
RS232
Inside NB05
Every 100 s the onboard TIC measures
  PPSout - PPSin (XO)
  RFin - PPSin
Only positive delays!

## Slide 19: NIST Travelling System
Inside NB05
5/10 MHz
1 PPS
RS232
RFin is amplified, power and frequency coarsely measured
PPSin and PPSout are re-triggered at 0.5V
PPSin, PPSout: TTL, 50 Ohms
RFin: 5 or 10 MHz, -3 to +5 dBm, 50 Ohm

## Slide 20: NIST Travelling System
Log file: file logged by LV app
controller provides:
  Receiver’s health
  RF power and frequency
  TIC measurements
  Delay line control
Inside NB05
5/10 MHz
1 PPS
RS232

## Slide 21: NIST Travelling System

## Slide 22: NIST Travelling System

## Slide 23
Receiver setup

## Slide 24: Calibration Trips Requirements
LABORATORY INFORMATION
Block diagram illustrating local setup, including the NIST travelling receiver
Information sheet (one for each local receiver)

## Slide 25: Calibration Trips Requirements
DATA AND FILE FORMATS
Local delays for NIST travelling receiver:
CABDLY (provided)
REFDLY = XP +XO
Measurements using provided cables, TIC and data logger
(beginning and end of visit)
Continuous logging after initial setup is completed
Delay-MMJDD.dat
TIC-NB05-MMJDD.dat
Local delays for LABk receiver:
CABDLY (from past measurements)
REFDLY = XP +XO
BEST: measured at beginning and end of visit
GOOD: data from latest measurement

## Slide 26: Calibration Trips Requirements
DATA AND FILE FORMATS
GPS data files for NIST travelling receiver
NB05_dow.bin
NB05doy0.YYO
NB05doy0.YYN
GPS data files for LABk receiver
Rinex
5 continuous days of GPS data are required as a minimum at each location
RINEX
Do not enforce clock corrections
CGGTTS
Ver. 2 or higher

## Slide 27: Calibration Trips Requirements
LOGISTICS
All data files need to be uploaded on a secure location provided by NIST
The duration of each calibration trip should be kept at a minimum
The NIST travelling system has a packing list and photographs of the packed enclosure to facilitate re-packing
INTERNATIONAL SHIPPING VIA ATA CARNET
Provides documentation for temporary import/export
Does not require additional custom brokerage by the country that recognize it
Requires reachable contact personnel at all locations indicated
The carnet must be with the package at all times (the shipping company must deliver it together with the package)

## Slide 28
BIPM Calibration Guidelines
THANK YOU!
for this opportunity
and
for participating in the SIM G2 campaigns
ftp://ftp2.bipm.org/pub/tai/publication/gnss-calibration/guidelines/

## Slide 29: UTC(k) Best Practices
Definition of Local Time Reference Point
PPS distribution amplifier
LOCAL
TIME REFERENCE POINT
USER
USER
USER
USER
The Local Time Reference Point must be accessible at all times!

## Slide 30: Local Delays: REFDLY
“REFDLY=(delay between UTC(k) and PPSin) + (delay between PPSin and RCVR’s internal ref. point)”
XO
When PPSout is available
PPSout is the best representation of the receiver’s time base
Can be measured using a PIVOT (measuring PPSout-UTC(k) and using XP)
Should be continually monitored (provides info about health of receiver’s time base)
PPSin
RFin
PPSout
PLL
Divide by X∙106
ENABLE

## Slide 31: Local Delays: REFDLY
“REFDLY=(delay between UTC(k) and PPSin) + (delay between PPSin and RCVR’s internal ref. point)”
XO
When PPSout is NOT available
XO ≠ 0
the receiver’s internal reference point has been defined in terms of a zero-crossing of RFin
can be measured using a PIVOT (measuring RFin-UTC(k) and using XP)
XO = 0
PPSin is accepted as the receiver’s internal reference point

## Slide 32: Delay measurements
Delay measurements using PPS signals and TIC are sensitive to:
  Pulse rise time (faster pulses are better)
  Cable losses and impedance
  TIC non-linearities and bias
Delay measurements using a PIVOT:
  Always measure positive delays
  Are immune to TIC biases and most of the issues listed above reducing systematic uncertainties (when cable lengths are similar)
If UTC(k) reference point and the receiver are far from each other, this delay should be measured using a Clock Trip.

## Slide 33: UTC(k) Best Practices
Local Dissemination
The delay of the PPS signal at each user should be regularly measured to monitor possible changes.
Distribution via coaxial cable requires measurements every year
Distribution via active systems (i.e. active oprical fiber) may require measurements over shorter intervals
Co-located users require TIC and a pivot point
Users in different rooms/buildings require a clock trip between the user and the local time reference point (TIC and a portable clock)

## Slide 34: NIST Travelling System

## Slide 35: Differential calibrations
Differential calibration of GPS link (NIST and CGGTTS definitions)
PPSin
rcvr time ref
(PPSout)
antenna
CABDLY1
t1
(INTDLY1)
RCVR1
LOCAL
TIME REFERENCE PLANE
UTC1
PDA
Antenna (INTDLY1)
REFDLY1
PPSin
rcvr time ref
(PPSout)
antenna
CABDLY2
t2
(INTDLY2)
RCVR2
LOCAL
TIME REFERENCE PLANE
UTC2
PDA
Antenna (INTDLY2)
REFDLY2
CV1T − CV2T = INTDLY1 – INTDLY2
Differential calibration of GNSS equipment
CABDLYT
tT
(INTDLYT)
RCVRT
REFDLYT
Antenna (INTDLYT)
REFSYS1 − REFSYST = INTDLY1 − INTDLYT
REFSYS2 − REFSYST = INTDLY2 − INTDLYT
RCVRT at ①
RCVRT at ②

## Slide 36: Absolute calibration
simulator