## Slide 1
Calibration of optical tachometers
Physics Metrology Division
Nelson Bahamón Cortés (nbahamon@inm.gov.co)
Andrés Montaño Rodriguez (wamontano@inm.gov.co)

## Slide 2
What is an optical tachometer?

## Slide 3
Optical Tachometer
Contactless type instrument that uses light beam (IR or laser) to measure rotational frequency, i.e: to count amount of engine´s shaft rotations per time unit.

## Slide 4
Technical Specifications Tachometer DT 2236B
Range | (2.5 to 99 999) RPM | Time base | Quartz cristal
Resolution (Experimental) | 0.01 RPM (2.5 to 999) RPM 0.1 RPM (1 000 to 9 999) RPM 1 RPM (> 9 999 RPM) | Sampling time | 1 second (over 6 RPM)
Display | 5 digits LCD | Measuring Distance | Max. 2 m
Accuracy | + (0.05% + 1 digit) | Measuring Angle | At less than 120°
Auto zero adjustment |  |  | 

## Slide 5
Quantity and Measure Unit
Rotational Frequency
[RPM]
On repetitive physical phenomenon, frequency is usually measured using revolutions per minute units and its symbol es r/min or RPM.
1 Hz
60 RPM
=
Frequency
[Hz]
On electrical signals, standard unit is Hertz, its symbol is Hz and it is defined as the number of events or cycles per second.

## Slide 6
Calibration System
Signal Generation System
  Waveform generator and counter, both disciplined to an standar signal traceable to SI unit.
  Light emitter diode to generate optical pulses.

## Slide 7
Calibration Procedure
Direct comparison calibration against frequency measurement of a frequency counter traceable to SI unit.
The time Interval between each pulse corresponds to the period of the signal frequency.

## Slide 8
Recommendations
Alignmet: keep DUT and LED front to front, you can use a metal holder to attach the DUT.
Measurement Time: take into account the period of the signal frequency.

## Slide 9
Method of measurement

## Slide 10
Mathematical model – estimation of uncertainty
Standard Deviation vs Allan Variance

## Slide 11
Mathematical model – estimation of uncertainty
Standard Deviation vs Allan Variance

## Slide 12
Modelo matemático
Desviación estándar vs Varianza de Allan

## Slide 13
Modelo matemático

## Slide 14
Central limit theorem
u1
u2
u3
u4
u5
u6
u7
u8
uC
Mathematical model – estimation of uncertainty

## Slide 15
u1
u2
u3
u4
u5
u6
u7
uC
uD
Mathematical model – estimation of uncertainty
Central limit theorem. Dominance

## Slide 16
Mathematical model:
Mathematical model – estimation of uncertainty

## Slide 17
Mathematical model:
Uncertainty:
Repeatability Resolution Standard
Mathematical model – estimation of uncertainty

## Slide 18
Component of uncertainty:
Repeatability:
Resolution:
Standard:
Mathematical model – estimation of uncertainty

## Slide 19
Component of uncertainty:
Repeatability:
Resolution:
Standard:
Mathematical model – estimation of uncertainty

## Slide 20
Component of uncertainty:
Repeatability:
Resolution:
Standard:
Mathematical model – estimation of uncertainty

## Slide 21
Combined uncertainty:
Mathematical model – estimation of uncertainty

## Slide 22
Mathematical model – estimation of uncertainty
Combined uncertainty:

## Slide 23
Mathematical model – estimation of uncertainty
Combined uncertainty:

## Slide 24
Mathematical model – estimation of uncertainty
Combined uncertainty:

## Slide 25
Expanded uncertainty
Mathematical model – estimation of uncertainty

## Slide 26
Dominance:
Mathematical model – estimation of uncertainty
Calibration Guide EURAMET cg-15 Versión 3.0 (02/2015)

## Slide 27
Welch-Satterwaite:
Mathematical model – estimation of uncertainty

## Slide 28
Welch-Satterwaite: GUM Anexo G
Mathematical model – estimation of uncertainty

## Slide 29
Welch-Satterwaite: GUM Anexo G
= 200
Mathematical model – estimation of uncertainty

## Slide 30
t distribution:
Mathematical model – estimation of uncertainty

## Slide 31
t distribution:
GUM Tabla G.2
Mathematical model

## Slide 32
Normal distribution:
GUM Tabla G.1
Mathematical model

## Slide 33
GUM Tabla G.1
Mathematical model
Normal distribution:

## Slide 34
Rectangular distribution:
GUM nota tabla G.1
Mathematical model

## Slide 35
Instituto Nacional de Metrología de Colombia - Av. Carrera 50 No 26 - 55 Int. 2 Bogotá, D.C. - Colombia Conmutador: (571) 2 54 22 22 / 5 88 02 46 E-mail: contacto@inm.gov.co Website: www.inm.gov.co
¡Gracias!